## Slide 1: Measurement of the magnet center with Single Stretched Wire system
Y. Arimoto, N. Ohuchi,
M. Masuwawa, R. Ueki, Y. Ohsawa, T. Kawamoto
and H. Iinuma

## Slide 2
in stage coordinate system

## Slide 3
Co-directional motion (finds nodal point (average axis))
Counter-directional motion (to find true axis)
Stage motion causes the wire to move in the B- field to generate change in flux, Φ, measured by the SSW data acquisition system.
Output wire ends points in stage coordinates (x, y)
  Input z positions from design
  Sag corrections
Measure roll angles, too.
Find magnets center positions (x, y)
Move wire ends till CO0 and CN0

## Slide 4
We care the stage levels ~ 0.17mrad
After earthquake, we need to arraign them again
SSW system delivered on May 12
Test bench, just across the street

## Slide 5: Repeatability check
Red 10 points
Blue 4 points
Repeat from the wire set

## Slide 6: To convert fix coordinate, we apply laser tracker survey
Survey data is consistent with design values and stable within few tens of microns

## Slide 7: Survey by Laser tracker

## Slide 8
We added 4 points for each stage to increase accuracy

## Slide 9: Measured magnet centers vs. designs: HER
Magnet center positions with several conditions (AC: 300K, ~50K, 4K and DC) are stable within 50~80m for x and y directions
AC 300K

## Slide 10: Measured magnet centers vs. designs :LER
AC 300K

## Slide 11: Corrector coil measurement

## Slide 12: Corrector coil measurements
QC1LE+A2
ハーモニックコイルの結果よりも10%多めに磁場が回転しています

## Slide 13
Roll Angle Summary
-2 mrad
-2 mrad
-15 mrad
-7 mrad

## Slide 14: Summary
From measurements so far：
Magnet center positions with several conditions (AC: 300K, ~50K, 4K and DC) are stable within 50~80m for x and y directions
Horizontal magnet centers of QC2LP and QC2LE differ ~1mm from the design.
  True?
  B1 corrector coil can correct
  How about QC1LP and QC2LP ?  we will measure soon.
Vertical positions differ 100～600 m compared to the design values
  True?
  Hard to correct especially for QC1LP
Roll angles also differ ~2mrad.
  A2 corrector coils are available
  Careful comparison with Harmonic coils’ dataと
Next:
We will resume SSW measurement in the next week.
Understanding (hidden) systematic uncertainties … use normal magnets
Measurements of QCSR in November…
seeing is believing

## Slide 15: Return is also in the pipe

## Slide 16
Axis Repeatability

## Slide 17
Strength measurements at full field
Need to understand variations – e.g. slight wire susceptibility may influence large step measurements in strong magnet (perhaps sag calibration measurements can be used to extrapolate this to “infinite tension”)

## Slide 19: QC1LE+A2
ハーモニックコイルの結果よりも10%多めに磁場が回転しています

## Slide 20: QC1LE
 |  |  |  |  |  |  |  |  |  |  | sag_correction | 
 |  |  | trackStageA | trackStageB | trackMagA | trackMagB | center | TrackSagMagA | TrackSagMagB | center | magA | magB
20160517_221336 | AC 4.5K | x | 167.617 | 12.08 | 65.70022453 | 50.10140086 | 57.9008127 | 65.71455733 | 50.08015574 | 57.89735654 | 0.014333 | -0.0212451
SURVEY0 |  | y | 0.512 | -0.509 | -0.15701784 | -0.259414048 | -0.208215944 | -0.091425432 | -0.323755047 | -0.20759024 | 0.065592 | -0.064341
20160518_002642 | DC | x | 167.637 | 12.06 | 65.69401424 | 50.09117896 | 57.8925966 | 65.71852526 | 50.07478121 | 57.89665323 | 0.024511 | -0.0163978
SURVEY0 |  | y | -3.808 | 1.321 | -0.447184623 | 0.067203379 | -0.189990622 | -0.475856962 | 0.142960117 | -0.166448423 | -0.02867 | 0.07575674
20160518_234111 | AC53K | x | 167.587 | 12.09 | 65.69643483 | 50.10162276 | 57.8990288 | 65.69814867 | 50.10102994 | 57.89958931 | 0.001714 | -0.0005928
SURVEY2_10mmstep |  | y | -0.178 | -0.269 | -0.237628426 | -0.246754827 | -0.242191627 | -0.266604528 | -0.213391969 | -0.239998249 | -0.02898 | 0.03336286
20160524_101959 | AC 4.7K | x | 167.487 | 12.21 | 65.74059147 | 50.16784322 | 57.95421734 | 65.74757512 | 50.15436334 | 57.95096923 | 0.006984 | -0.0134799
SURVEY3 |  | y | 1.362 | -0.729 | -0.008143294 | -0.217849925 | -0.112996609 | -0.007694291 | -0.20179481 | -0.10474455 | 0.000449 | 0.01605512
20160524_124446 | DC | x | 166.667 | 12.55 | 65.68069009 | 50.22427838 | 57.95248423 | 65.63983751 | 50.26843836 | 57.95413793 | -0.04085 | 0.04415998
SURVEY3 |  | y | 0.012 | -0.209 | -0.132811893 | -0.154976008 | -0.143893951 | -0.166476934 | -0.117781314 | -0.142129124 | -0.03367 | 0.03719469
20160512_202305 | AC room | x | 167.527 | 12.36 | 65.85266978 | 50.29095345 | 58.07181162 | 65.86419547 | 50.27400258 | 58.06909902 | 0.011526 | -0.0169509
SurveyWarm1 |  | y | -1.978 | 0.431 | -0.399484842 | -0.157885954 | -0.278685398 | -0.381126052 | -0.192837735 | -0.286981893 | 0.018359 | -0.0349518
 |  |  |  |  |  |  |  |  | ave_x | -57.94 | design_x | -57.95
 |  |  |  |  |  |  |  |  | ave_y | -0.19 | design_y | 0

## Slide 21: QC2LE
 |  |  | trackStageA | trackStageB | trackMagA | trackMagB | center | TrackSagMagA | TrackSagMagB | center | magA | magB | 
20160518_024910 | DC | x | 169.827 | 9.34 | 123.884311 | 100.712315 | 112.2983 | 123.82473 | 100.76752 | 112.2961 | -0.05958 | 0.055203 | 
SURVEY1 |  | y | -0.898 | 0.591 | -0.4717433 | -0.2567533 | -0.36425 | -0.4641617 | -0.259269 | -0.36172 | 0.007582 | -0.00252 | 
20160518_035039 | AC 4.5K | x | 169.307 | 10.28 | 123.782266 | 100.821073 | 112.3017 | 123.77206 | 100.81551 | 112.2938 | -0.0102 | -0.00557 | 
SURVEY1 |  | y | -0.748 | 0.351 | -0.4333888 | -0.2747091 | -0.35405 | -0.4425508 | -0.256245 | -0.3494 | -0.00916 | 0.018464 | 
20160518_195218 | AC | x | 169.287 | 10.37 | 123.793755 | 100.848445 | 112.3211 | 123.80034 | 100.81964 | 112.31 | 0.006588 | -0.02881 | 
SURVEY2 |  | y | -0.638 | 0.181 | -0.4035445 | -0.2852928 | -0.34442 | -0.3881773 | -0.291196 | -0.33969 | 0.015367 | -0.0059 | 
20160519_043749 | DC | x | 169.957 | 9.1 | 123.908391 | 100.682972 | 112.2957 | 123.90066 | 100.70762 | 112.3041 | -0.00773 | 0.024644 | 
5mmstep+SURVEY2 |  | y | -1.008 | 0.881 | -0.4672351 | -0.1944909 | -0.33086 | -0.5701009 | -0.22443 | -0.39727 | -0.10287 | -0.02994 | 
20160525_102933 | AC | x | 169.307 | 10.24 | 123.770815 | 100.803846 | 112.2873 | 123.77013 | 100.80837 | 112.2892 | -0.00069 | 0.004522 | 1/f2
SURVEY4 |  | y | -0.868 | 0.671 | -0.4274297 | -0.2052205 | -0.31633 | -0.4146487 | -0.222609 | -0.31863 | 0.012781 | -0.01739 | 1/f2
20160525_140642 | DC | x | 169.687 | 9.58 | 123.853094 | 100.735964 | 112.2945 | 123.85557 | 100.73068 | 112.2931 | 0.002481 | -0.00528 | 1/T
SURVEY4 |  | y | 0.542 | -2.139 | -0.2254911 | -0.6125887 | -0.41904 | -0.1791846 | -0.66073 | -0.41996 | 0.046307 | -0.04814 | 1/T
20160513_001236 | AC_room | x | 169.217 | 10.52 | 123.786735 | 100.873189 | 112.33 | 123.81087 | 100.84881 | 112.3298 | 0.024137 | -0.02438 | 
surveyWarm2 |  | y | -0.208 | -0.449 | -0.2769912 | -0.3117881 | -0.29439 | -0.3146003 | -0.305977 | -0.31029 | -0.03761 | 0.005811 | 
 |  |  |  |  |  |  |  |  | ave_x | -112.30 | design_x | -111.32 | 
 |  |  |  |  |  |  |  |  | ave_y | -0.36 | design_y | 0 | 

## Slide 22: QC1LP & QC2LP
 |  |  | trackStageA | trackStageB | trackMagA | trackMagB | center | TrackSagMagA | TrackSagMagB | center | magA | magB | 
20160516_045401 | AC | x | -166.783 | -13.79 | -46.312957 | -32.572255 | -39.4426 | -46.3466782 | -32.53125995 | -39.438969 | -0.033721 | 0.040995055 | 
surveyA |  | y | -2.188 | -1.579 | -1.70846 | -1.65376416 | -1.68111 | -1.70778584 | -1.654880837 | -1.6813333 | 0.0006742 | -0.001116676 | 
20160516_222636 | DC | x | -171.243 | -12.84 | -46.513005 | -32.2864161 | -39.3997 | -46.554758 | -32.26478877 | -39.409773 | -0.041753 | 0.021627337 | 
surveyB |  | y | -4.138 | -1.159 | -1.7922701 | -1.52471829 | -1.65849 | -1.80819186 | -1.505194702 | -1.6566933 | -0.015922 | 0.019523584 | 
20160513_025418 | AC_room | x | -168.833 | -13.32 | -46.378654 | -30.361624 | -38.3701 | -46.3225859 | -30.42501199 | -38.373799 | 0.0560682 | -0.063388043 | 
surveyWarm3 |  | y | -9.578 | -0.889 | -2.7360909 | -1.95570902 | -2.3459 | -2.61632632 | -2.08254121 | -2.3494338 | 0.1197646 | -0.126832186 | sagBig?
 |  |  |  |  |  |  |  |  | ave_x | 39.07 | design_x | 38.91 | 
 |  |  |  |  |  |  |  |  | ave_y | -1.90 | design_y | -1.5 | 
 |  |  | trackStageA | trackStageB | trackMagA | trackMagB | center | TrackSagMagA | TrackSagMagB | center | magA | magB
20160516_024430 | AC(~4K) | x | -167.813 | -14.61 | -89.7101568 | -72.80354746 | -81.25685213 | -89.69582928 | -72.8192085 | -81.25751889 | 0.014328 | -0.01566
surveyD |  | y | -2.288 | -2.079 | -2.181451863 | -2.158387815 | -2.169919839 | -2.179956122 | -2.155532459 | -2.16774429 | 0.001496 | 0.002855
20160517_022234 | DC | x | -172.913 | -10.79 | -90.26274348 | -72.37177383 | -81.31725866 | -90.21191545 | -72.43588234 | -81.32389889 | 0.050828 | -0.06411
suevryC |  | y | -6.538 | 1.201 | -2.592660134 | -1.738628231 | -2.165644182 | -2.519728924 | -1.950059456 | -2.23489419 | 0.072931 | -0.21143
20160513_043512 | AC_room | x | -168.323 | -14.36 | -89.83270902 | -72.84223043 | -81.33746973 | -89.78941716 | -72.88746965 | -81.3384434 | 0.043292 | -0.04524
surveyWarm4 |  | y | -2.308 | -2.029 | -2.165765884 | -2.134977035 | -2.15037146 | -2.170340417 | -2.159442026 | -2.164891222 | -0.00457 | -0.02446
 |  |  |  |  |  |  |  |  | ave_x | 81.31 | design_x | 80.11
 |  |  |  |  |  |  |  |  | ave_y | -2.19 | design_y | -1.5

## Slide 23: QC2LE+A1, B1 and A2
 | QC2LE_correction coils |  |  | dx | dy
 | SSW_R_20160525_200006_DC_SO | 2016/05/25 08:00:10.618 PM | QC2LE_main | 0 | 0
A1 | SSW_R_20160525_200348_DC_SO | 2016/05/25 08:03:51.900 PM | QC2LE+A1_+60A | 0.0140205 | 3.1974287
 | SSW_R_20160525_200555_DC_SO | 05:58.8 | QC2LE+A1_+60A | 0.0162405 | 3.1938276
 | SSW_R_20160525_201033_DC_SO | 2016/05/25 08:10:37.272 PM | QC2LE+A1_-60A | 0.0811598 | -2.960667
 | SSW_R_20160525_201222_DC_SO | 2016/05/25 08:12:26.095 PM | QC2LE+A1_-60A | 0.0795518 | -2.960187
 | SSW_R_20160525_201514_DC_SO | 2016/05/25 08:15:18.648 PM | QC2LE+A1_-60A | 0.0788113 | -2.957992
B1 | SSW_R_20160525_204744_DC_SO | 2016/05/25 08:47:48.893 PM | QC2LE+B1_+60A | 3.4432189 | 0.0451268
 | SSW_R_20160525_205039_DC_SO | 2016/05/25 08:50:43.432 PM | QC2LE+B1_+60A | 3.450505 | 0.0439505
 | SSW_R_20160525_205524_DC_SO | 2016/05/25 08:55:28.708 PM | QC2LE+B1_+60A | 3.4449356 | 0.0481159
 | SSW_R_20160525_212356_DC_SO | 2016/05/25 09:24:00.510 PM | QC2LE+B1_+30A | 1.6387746 | 0.0264326
 | SSW_R_20160525_212816_DC_SO | 2016/05/25 09:28:20.002 PM | QC2LE+B1_-30A | -1.903114 | 0.0035128
 |  |  |  | Angle.avg | Angle.sdev
 | SSW_R_20160525_214658_DC_RA | 2016/05/25 09:47:02.568 PM | QC2LE | -2.284629 | 0.1364434
A2 | SSW_R_20160525_215436_DC_RA | 2016/05/25 09:54:40.081 PM | QC2LE+A2_+30A | 13.45524 | 0.119109
Harmonic coilの結果との比較も準備します。
ヒステリシス5%くらい

## Slide 24: QC1LE+A1, B1 and A2
 | QC1LE_correction coils |  |  | dx | dy
 | SSW_R_20160526_144713_DC_PY | 2016/05/26 02:47:17.466 PM | QC1LE | 0 | 0
B1 | SSW_R_20160526_150302_DC_SO | 2016/05/26 03:03:05.936 PM | QC1LE+B1_+30A | 1.056283 | -0.01793
 | SSW_R_20160526_151312_DC_SO | 2016/05/26 03:13:16.228 PM | QC1LE+B1_-30A | -1.05438 | 0.006157
 | SSW_R_20160526_151629_DC_SO | 2016/05/26 03:16:33.615 PM | QC1LE+B1_-60A | -2.1083 | 0.013085
A1 | SSW_R_20160526_152959_DC_SO | 2016/05/26 03:30:03.461 PM | QC1LE+A1_+30A | -0.01298 | 0.622072
 | SSW_R_20160526_153649_DC_SO | 2016/05/26 03:36:53.593 PM | QC1LE+A1_-30A | -0.01758 | -0.62879
 | roll |  |  | Corrected.Angle.avg | Corrected.Angle.sdev
 | SSW_R_20160526_160723_DC_RA | 2016/05/26 04:07:27.939 PM | QC1LE | -2.20886 | 0.078277
A2 | SSW_R_20160526_163254_DC_RA | 2016/05/26 04:32:58.000 PM | QC1LE+A2_30A | 7.076841 | 0.060697
 | SSW_R_20160526_164510_DC_RA | 2016/05/26 04:45:14.000 PM | QC1LE+A2_-30A | -11.4009 | 0.061995
 | SSW_R_20160526_165800_DC_RA | 2016/05/26 04:58:03.922 PM | QC1LE+A2_-60A | -21.2021 | 0.072452
 | SSW_R_20160526_171628_DC_RA | 2016/05/26 05:16:32.788 PM | QC1LE | -2.93344 | 0.078084
A2 | SSW_R_20160526_172655_DC_RA | 2016/05/26 05:26:58.874 PM | QC1LE+A2_+60A | 16.56536 | 0.082787
Harmonic coilの結果との比較も準備します。