## Slide 1
QCS
The 27th KEKB Accelerator Review Committee 2024/MAR/26th
Toshiyuki OKI on behalf of the SuperKEKB QCS Group

## Slide 2
Outline
  Main work carried out during LS1
    (1) Modifications of the head of QCS cryostat to reduce background
    (2) Repair of vacuum leakage at the QCSR service cryostat
  Operational status (cooling, magnet power supply)
  Plan for anti-ageing of old system
    (1) Overhaul of electric motor for helium compressor
    (2) Refrigeration control system renewal
  Considerations
    (1) Additional heat exchanger
    (2) Cooling margin of refrigerator

## Slide 3
Modification work around IR to reduce background
  Before LS1, the beam background production points were evaluated by the Belle group.
  During LS1, modification work was carried out around the IR to reduce background: the Belle group added several types of shields, and we, QCS group, together with the vacuum group, modified the front part of the QCS cryostat.
  K. Nakamura. Mar 4, 2021
  SuperKEKB IR meeting
  made by heavy metal

## Slide 4
QCSL cryostat: head plate replacement
  The tungsten head plate at the front part of the QCSL cryostat was replaced with a stainless steel one.
  For this replacement work, the beam pipes had to be shaked across the beam axis, so that the flanges between the cryostat and beam pipes were removed. In preparation for that, the cryostat front section was supported on a stand while the cryostat middle section was opened to allow access to the beam pipe support screws and BPM cables. Displacement measurements were carried out before and after the work.
head plate

## Slide 5
QCSR cryostat: front cap with head plate replacement
  For the QCSR, the tungsten head plate was also replaced with a stainless steel one.
  Furthermore, the front cap of the cryostat has been redesigned smaller along with the inner heat-insulating layers, so that the cable space for the VXD was enlarged.
  The QCSR cryostat has service ports from which the beam pipes can be accessed.
front cap

## Slide 6
Vacuum leak repair work on the QCSR service cryostat
  Before LS1, a vacuum leak had been found in the upper and lower connection of the QCSR service cryostat.
  During LS1, it was investigated for identification of the leakage part and repaired.
  Fibrous material
  Milling step
  Service
  cryostat
  Polished by whetstone

## Slide 7
QCSR leak rate and vacuum pressure after repair
  The leakage rate of the QCSR cryostat was high before the repair work, but improved to the same level as the QCSL after the repair.
  A cooling test in December 2023 confirmed that the vacuum pressure could be reached at the same level as the QCSL both at room temperature and at low temperatures.

## Slide 8
Current status of QCS: cooling
  According to the cooling characteristic curve at the start of cooling, the QCS has been cooled well without any negative effects due to the cryostat modification work during LS1.

## Slide 9
Current status of QCS: operation
  From the start of 2024a run to the present, the QCS magnet power supply has operated stably without failure with a current stability of a few ppm/week.
  As reported in the previous ARC, the QC1 magnet initialisation pattern for ramp-up is also used in 2024a run to mitigate the flux creep in superconducting magnets.
  Ramp-up pattern for QC1P
  Current stability for a week (5ppm/div)
  QC1LP
  QC1RP
  QC2LP
  QC2RP
  QC1LE
  QC1RE
  QC2LE
  QC2RE
  ESL
  ESR1
  ESR23
  5ppm

## Slide 10
Overhaul/renewal of deteriorated system: motors for He compressor
  There are three electric motors for the helium compressor, one each for the QCSL/QCSR and one for a spare. These are over 30 years old and beyond their usual lifetime, so an overhaul is planned.
  To allow overhauls to be carried out during short shutdown periods, the plan is to overhaul the three motors in sequence, one each year; one was done last year and the other will be done this year.

## Slide 11
Overhaul work for the QCSL He compressor motor in 2023 (1/2)
  Overhaul of a motor for QCSL He compressor was carried out at a factory in 2023.
  An electrical test was carried out just before the overhaul work.
  The results showed that the worst-case remaining lifetime was only about one year. The need for overhaul for the remaining motors is strongly recognised.

## Slide 12
Overhaul work for the QCSL He compressor motor in 2023 (2/2)
  At the factory, the bearings were replaced, coils were rewound, and various adjustments and tests were carried out.
  The overhauled motor has been applied to the QCSL and is currently operating without problems. Overhaul for QCSR motor is planned in summer 2024.

## Slide 13
Overhaul/renewal of deteriorated system: refrigerator control system
  The control system controls the refrigerator from Nikko control room to Tsukuba area (consisting of more than 300 control loops).
  The system has been in use since the KEKB era and is no longer supported by the fabrication maker. Some of them have no replacement parts.
  The control system needs to be updated to operate for the next 10 years or so.
  The most important thing to keep in mind is not to destroy the existing refrigerator system (the refrigerator system cannot be replaced. It is too expensive and impractical to renew them.).
  Nikko control room

## Slide 14
Renewal Plan
  There are two main options for the update plan.
      Install successor system of the current system: All design (including sequence design) is done by the manufacturer.
      Build a new system by ourselves: All design is done by ourselves, where the system will be based on PLC. The cooling control scheme in the software is the same as the present system.
  Current system
  PLC based home-made system

## Slide 15
Pros. & Cons.
 | (A) Successor | (B) Home-made
Cost | 790k USD (Belle+QCS) | Low (R&D cost is unknown)
Manpower | a few | Large
Compatibility (system, operation) | Good | Poor
Risk to destroy a refrigerator or stop system | Low | Need to debug with spare system
  Note: The QCS and Belle refrigeration control systems share information and infrastructure facilities with each other. It is therefore preferable to renew them at the same time in the same system. In particular, Belle's cryogenic group has commented that they want to avoid the uncertainty risk from the development required to introduce a new environment.
  Coupled Belle & QCS system

## Slide 16
Reducing the use of LN2: consideration of additional heat exchangers
  To reduce the LN2 quantity consumed for the cryogenic system (ECO-system), we have started to discuss about the installation of the additional heat exchanger.
  The amount of the LN2 quantity is 3500 L/day for the QCS-R and QCS-L helium cryogenic systems. The cost of the LN2 per year is about 40 MJ\ (267 kUS$).
  By the additional heat exchanger, the amount of the LN2 consumption can be reduced to be smaller than the half.
  For the installation, the cost is estimated to be 17 MJ\ × 2 units. The work period in the practical setting is about 2 months (and for construction of the heat exchanger, 9 months).
  There is no risk for the additional heat exchanger, and the exchanger enhance the stability of the helium refrigerator.

## Slide 17
Reducing the use of LN2: Thermal condition of QCS cryogenic system
Magnet Cryostat
Service Cryostat
He Refrigerator
LHe Subcooler
1st-H.X.
2nd-H.X.
4th-H.X.
3rd-H.X.
Cold Box
Additional H.X.
Magnets: 4.5 K
LHe: 4.35 K, 0.16 MPa
Present thermal condition at the 1st H.X.
1st H.X.
GN2
High pressure GHe
Low pressure GHe
1.7 MPa
298 K
44.58 g/s
1.7 MPa
86.2 K
0.12 MPa
274 K
16.3 g/s
0.12 MPa
82 K
0.111 MPa
282 K
43.4 g/s
0.111 MPa
79 K
By adding the Heat Exchanger,
1st H.X.
GN2
High pressure GHe
Low pressure GHe
1.7 MPa
298 K
44.58 g/s
1.7 MPa
86.2 K
0.12 MPa
274 K
3.75 g/s
0.12 MPa
82 K
0.111 MPa
290 K
43.4 g/s
0.111 MPa
79 K
Additional H.X.

## Slide 18
Reducing the use of LN2: Thermal condition of QCS cryogenic system
1st Heat Exchanger between high and low-pressure helium gases and nitrogen gas .
The additional heat exchanger is expected to be one third of the 1st H.X..
QCS-R Helium refrigerator and sub-cooler
Additional H.X.
Location of the additional heat exchanger

## Slide 19
Cooling-power margin of the refrigerators
  We have begun to investigate the margin of cooling capacity in preparation for future upgrade plans.
  How to estimate current cooling-power margin:
    The liquid level in the receiver is controlled to prevent liquid helium from overflowing by using a heater to evaporate the extra helium produced.
    Heater power is as the refrigerator power margin.
  Pheater = Pmargin
  Decrease heater power
  Increase heater power
  Liquid Helium
  from Liquefier to Magnets
  Liquid Helium
  from Liquefier to Magnets

## Slide 20
History of the heater power of QCSR ( margin of the refrigerator )
  The heater power varied and exhibited 0 watt, sometimes. The variation caused by change of operation conditions such as the liquid nitrogen filling. A heater power of 0 W indicates that the heat load exceeds the liquid helium supply.
  Even if the control is well adjusted and the heater power can be smoothed, the margin is only ~20 W for a cooling power of 250 W.
  We need careful consideration to maintain the margin of the cooling power.
  80
  40
  0
  Heater power (W)
  Heater power of QCSR, 7 days in June 2022
  The average is ~20 W.

## Slide 21
Summary
  During LS1, the head of QCSL/QCSR cryostats were modified to reduce background. Vacuum leakage at the QCSR service cryostat was also repaired.
  QCS magnet power supply and cooling system have operated stably from the start of 2024a run to the present.
  Plans to overhaul or update ageing systems: electric motors for helium compressors are planned to be overhauled annually. Consideration has started to be given to renewing the refrigeration control system.
  Consideration for the additional heat exchangers to reduce liquid nitrogen and the cooling margin of refrigerator has begun.

## Slide 22
Backup

## Slide 23
SuperKEKB IR
Tsukuba Exp. hall
Central control room for refrigerators
Nikko Exp. Hall
SuperKEKB Accelerator central room
Belle II cryogenic system
QCS-R cryogenic system
QCS-L cryogenic system
To display some VIP parameters
STF
cERL
KEKB SRF cavities
Control systems and operations of cryogenic systems
Operations (QCS-L&R+BelleII):
KEK staffs: on-call
3 operators/shift, 3 shifts/day
Same control room in the Nikko Hall with cryogenic system of SuperKEKB SRF cavities, STF and cERL.

## Slide 24
QCS systems of SC magnets, cryostats and cryogenic systems
Subcooled cooling scheme (stable, temperatures lower than saturated LHe)
Mixed loads of refrigeration and liquefaction (250 W in total, to cool 51 leads in QCS-L and 59 in QCS-R, 40 L/hour)
Total operation time: ~ 30 khours
QCS quench events: 66 (48 due to beams), recently due to earthquake
48/66

## Slide 25
QCS systems of SC magnets, cryostats and cryogenic systems
Underground devices (refrigerators, subcoolers, cryostats, instrumentations and quench detectors)
Ground devices (compressors, helium tanks, liquid nitrogen tanks)